## Slide 1: Validation

## Slide 2: Validation Data
4 Measured aruco trials for the corresponding validation path
Compare measured vs ground truth:
  Step 1: Load 24-trial derived registration matrix ( aruco kuka)
  Step 2: Transform all measured aruco points to kuka reference frame
  Step 3: calculate path lengths (mm) for each sequential trajectory
  Step 4: Get respective trajectory RMSE

## Slide 3: Calibration and Validation JSON
Original JSON FILE data (m)
Plotted in millimeters!

## Slide 5: Trajectories
Compare ground truth trajectories (Validation JSON) to the measured aruco ( transformed to kuka reference frame) coordinates.
4 Validation trials were used
Trajectories
0 1
1 2
23

## Slide 6: Trajectories

## Slide 7: Trajectory RMSE ** all on validation data
High RMS error in Z…. Need to investigate

## Slide 8: Look into Positional accuracy

## Slide 9: Positions in Kuka reference frame
Note that validation coordinates are outside the volume of the the calibration positions

## Slide 10: Validation GT vs Aruco Derived Positional Measurements
Use Transformation matrix: aruco kuka

## Slide 11: Positional RMSE
Point 0
Point 1
Point 2
Point 3
Points 2 and points 3 have the largest RMSE erros

## Slide 12: Hypothesis
Larger error may be due to calibration volume does not cover these points

## Slide 13: Check RMSE for Aruco measurements used for repeatability study
Recall that in the calibration study we used 5 points from the 29-calibration trial acquisition ( 24 for registration matrix + 5 for quality assessment)
 Test hypothesis. If so, we expect better positional accuracy for point within the calibration volume space

## Slide 14: See all together:
Using Validation data+JSON data
Using repeatability data + calibration JSON
Takeaways: Positional accuracy is better when points are located within the calibration volume space. Therefore, for future application this would need to be optimized for the working space.